# Studer A810 Calibration Quick Reference
The complete instructions are in the Service Manual section 4.2
##### important definitions
reference level = (magnetic) TAPE level	185, 250, 320, … nWb/m
operating level = voltage level on the INPUT/OUTPUT	studio: +4 dBu = 1,228 V = 0 VU aka linelevel
a tape recorded with reference level should produce operating level in the output in play mode (!)
##### the procedure
The calibrating steps should always be performed in the following order, and each section repeated for ch. 2.
#### 1 GENERAL PREPARATORY STEPS: 4.2.1.6
Clean the tape path. Set all switches on the PERIPHERY CONTROLLER to the correct setting 4.2.9.2
###### check the output level
- feed ch. 1 line input with 1kHz @ operating level
- connect mV-meter line to the output of ch. 1 and load with 600Ω
- press INP and release UNCAL
- adjust ouput level to operating level with the line amplifier trimmer accessible from the front.
###### check the VU-meter
- VU indication:	Operating level should result in a reading of 0 VU
- PPM indication:	Operating level should result in a reading of -6 (0 for reference level).
- correct the VU-meter reading with the trimmer potentiometer located on the back of the VU-meter amplifier.
#### 2 REPRODUCE ADJUSTMENTS: 4.2.2
set code switch 8 on: JS 8 = 1 to enable programming
###### preparatory steps 4.2.2.1
- set tape type, equalization and tape speed as desired
- connect mV-meter line to the output of ch. 1 and load with 600Ω
- press SAFE and REP, release UNCAL
- spool reference tape to ‘reference level’ section (0 dB)
###### reproduce level adjustment 4.2.2.2
- press CH 1 and LEVEL REPRO
- start recorder in PLAY
- read out reproduce level on the mV-meter
- adjust to desired line level by pressing UP or DOWN and press STORE
###### azimuth alignment of the reproduce head 4.2.2.3
- spool reference tape to ‘azimuth alignment section’ (-10 dB)
- always adjust to maximum level first and then to minimum phase difference! See SM
- rewind tape to ‘reference level section’, check level and repeat reproduce level adjustment if necessary!
###### frequency response adjustment 4.2.2.4
- spool to ‘frequency section’ of the tape (-20 dB)
- select the corresponding tone:
  16 kHz@30 ips, 14 kHz@15 ips, 12.5 kHz@7.5 ips, 8 kHz@3.75 IPS, all at -20 dB
- press CH 1 and TREBLE REPRO
- start the recorder in PLAY mode
- adjust to optimum frequency response by pressing UP or DOWN and press STORE
- rewind calibration tape to the 63 Hz tone (-20 dB)
- press CH 1 and BASS REPRO
- start the recorder in PLAY mode
- adjust for optimum frequency response by pressing UP or DOWN and press STORE
#### 3 RECORDING ADJUSTMENT: 4.2.3
tape speed 7.5 or 15 ips (or preferred studio speed). Use new (or practically new) unrecorded tape
###### preparatory steps 4.2.3.1
- connect 1 kHz at line level to input and mV-meter to output
- press READY and REP, release UNCAL
###### record level adjustment 4.2.3.2
- press CH1 and LEVEL RECORD on input keyboard
- put recorder in record mode (PvdM addition)
- read output level on mV-meter and press UP or DOWN key for adjusting to line level.
- press STORE
###### azimuth alignment of the record head 4.2.3.3
- set generator to 10 kHz and decrease level by 20 dB. Connect millivoltmeter to line output channel
- start machine in RECORD mode
- correct azimuth by turning screw. If severe adjustment was necessary, repeat preparatory steps
###### bias adjustment 4.2.3.4
- set generator to 10 kHz and level to 20 dB below line level. Connect millivoltmeter to line output
- press CH 1 and BIAS on input keyboard
- start machine in RECORD mode
- press DOWN until MIN light flashes
- search maximum output voltage with UP and write down this value
- continue with UP key until delta dB is reached. See table at page 4/56 (pg. 187 PDF). Press STORE
###### record level adjustments 4.2.3.6
- set generator to 1 kHz and operating level. Connect millivoltmeter to line output.
- press CH1 and LEVEL RECORD on input keyboard
- start machine in RECORD mode
- adjust to line level by pressing UP or DOWN and press STORE
###### frequency response adjustment 4.2.3.7
- set generator to line level -20 dB. Connect mV-meter to line output
- press CH1 and TREBLE RECORD on input keyboard
- start machine in RECORD mode
- adjust to optimum frequency response (upward of 1 kHz) by pressing UP or DOWN. Press STORE
- press BASS REPRO (PvdM verified)
- adjust to optimum frequency response (up to 1 kHz) by pressing UP or DOWN. Press STORE
###### Cross-talk adjustment 4.2.3.8 (!)
#### 4 SYNC REPRODUCTION ADJUSTMENT 4.2.4
not supported for 3.75 IPS; all parameters should be left at 00
###### preparatory steps 4.2.4.1
- Set jumper on REPRODUCE AMPLIFIER to narrow or wide, see 4.2.9.4
- connect mV-meter to line output
- select tape speed, tape type, and equalization
- press SAFE and SYNC buttons, release all UNCAL buttons
- mount calibration tape to REFERENCE LEVEL section. (0 dB)
###### sync reproduce level adjustment 4.2.4.2
- press CH1 and LEVEL REPRO on input keyboard
- start recorder in PLAY mode
- read out sync reproduce level and adjust to the desired line level (UP, DOWN) and press STORE
###### frequency response adjustments 4.2.4.3
- spool reference tape to the frequency responce section (-10dB) ← not correct (PvdM)
- press CH1 and TREBLE REPRO
- start recorder in PLAY mode
- play tone – 7.5 IPS: 8kHz, 15 IPS & 30 IPS: 12.5 kHz
- adjust for optimum frequency response by pressing UP or DOWN. Press STORE
- rewind reference tape to 63 Hz tone (-10 dB) ← again not correct (PvdM)
- press BASS REPRO (PvdM verified)
- start recorder in PLAY mode
- adjust to optimum frequency response by pressing UP or DOWN. Press STORE
DONE!
Now download and fill in the calibration table at https://www.reeltoreel.nl/studer/PvdM-miscellaneous/A810/A810%20calibrationtable/ for future reference.
##### eigen aantekeningen PvdM:

B67	operating level 0 VU = 0,775 V	= +0 dBu
A810	operating level 0 VU = 1,228 V	= +4 dBu

dus: 250 nWb/m tape op A810 geeft op	tape A: 0 dB tape-level weer als -1½ VU	~ 1,025 V	= +2½ dB
(welke snelheid??)	tape B: 0 dB tape-level weer als -3 VU	~ 880 mV	= +1.1 dB